## Weekly checksheet
AUCHENCORTH WEEKLY CHECKSHEET
DATE: | 2026-05-27
OPERATOR: | DH SL
WEATHER: | warm, slightly overcast
POLLUTION OBS: | Hen manure in gavins field - pungent…
REM 1
COMPUTER CHECK |  | LOGGERNET: ICOS_Periph
Updates installed | N | Voltage | 14.13
Updates required | N | TA and RH sensible
Logging to the S drive | Y | LOGGERNET: ICOS_SoilMet1_BkMet
NTP server check for 192.168.0.187 - record offset | -0.174 | Voltage | 13.87
NTP server check for 192.168.0.187 - record jitter | 0.028 | Any NANs/obvious issues | N
LOGGERNET: LOGGER (METMAST):
cnr4_fan_freq not 0 | Y | LOGGERNET: ICOS_SoilMet2
Any sensors reporting INF, Nan, -99999: | N | Voltage | 13.62
Battery Voltage ≥ 12V (number) | 13.31 | Any NANs/obvious issues | N
LOGGERNET: GASES: |  | Labview: EC_GP.vi
O3_CEH (ppb) | 10.5 | Wind Speed (ms-1) | 2.59
Offset to O3 analyser (ppb) | 0.2 | Sonic Data Updating | Y
NO_Blue | 0.41
Offset to Bluelight (ppb) | 0.01
NO2_Blue | 1.45
Offset to Bluelight (ppb) | 0.15
Voltage | 12.77
LabView: Biral VPF750: Present weather sensor
LOGGERNET: TOWER_T_Wind: |  | Data updating | Y
Battery Voltage ≥ 12V | 12.27 | S drive file upto date (PWS_VPF750_AU....dat) | Y
Thermocouples working (TC1-5) | Y
Wind speed sensors working (WS1-5) | Y | Phenocam
|  | Check Phenocam images downloading | Y
ICOS: LI-7x00: |  | Transfer Phenocam images to Server | Y
Data updating, logging | Y
USB Free Space >2GB | Y
File transfer running in system tray | Y
ANALYSER CHECKS
NO/NOx (blue light) - T200UP |  | LICOR
Alarm | N | LICOR N2 rotameter (is it just above 0) | y
Rx cell Temp | 40 | LICOR Pump Running and line connected to pump | y
Rx Cell pressure (inHg) | 2.5 | LICOR: Cell T (value) | 35.47
Sample pressure (in Hg) | 28.3 | LICOR: Cell P (value) | 25.14
Sample flow (cc/min) | 1157 | LICOR Calibration cylinder (Bar) | 60
O3 Flow (cc/min) | 88 | LICOR N2 Cylinder Pressure (Bar) | 160
Conv temp | 202.4
TEKRAN
O3 - T400 |  | Green Tick present on laptop loging software | y
Alarm | N | Data upto date | y
Temperature - O3 Gen Lamp | 48 | UV Lamp Indicator on analyser is white? | y
Temperature - Photo Lamp | 58 | Temperature in box 18C+/-4C (max hasn't exceeded 25C) | y
Sample pressure (in Hg) | 26.6 | Tekran Ar (High) Pressure (psi) | 1500
Sample flow (cc/min) | 861 | Air Con Working (TEKRAN) | y
O3 ref (mV) | 3967.2
Gas Chromatograph
Calibrator - T700 |  | GC H2 Generator Topped Up | y
Alarm | n | GC Compressor Bottle emptied | y
NO2 cylinder pressure (PSI) (Back wall) | 1350
Green lights, dew point and power on zero air generator | y | Wet only DWOC
|  | New Wet-Only Samples Exchanged | y
AURN Ozone |  | New Wet-Only Data Downloaded | y
AURN Ozone OK | y | New Wet-Only Changeover time verified in bottle.dat file | y
MARGA |  | BAM and EQ PARTISOL CHECKS
MARGA waste closed | y | BAM radiation sources present | y
MARGA inlet fan running | y | Tape - plenty left? | y
MARGA Solutions Exchanged | y | Partisol - on and filters avaliable? | y
MARGA suppressor/regenerant >1 inch left | y | FIDAS present and light on front? | y
Aethalometer
Aethalometer Green tick | y | eDNA partisol - wear gloves/mask!
Aethalometer Tape Left - open front and check | 72 | Sampler working OK | y
Aethalometer Flow - is it 5 LPM? | y | Filter collected & record sheet filled in | y
MET CHECKS - for ICOS sensor issues please report date, time of issues same day to Karen/Sarah
jNO2
Met Tipping Bucket- check | y | jNO2 Sensors/domes clear of debris / ice | na
ICOS tipping bucket - check | y | jNO2 OceanView running, data logging | na
Pluvio - check for something massive (dead bird size and above) | y | jNO2 data up to date on S:\JNO2\jNO2 data 2024 (two files: one for each instrument QEP043831 and QEP043912) | na
Main met mast |  | ICOS soil back up mast (date, time, issue must be reported same day if issue)
Radiation sensors - check clean (main met and radiation mast) | y | Radiation sensors - check clean | y
Temperature and RH shield -check clean | y | Check TRH aspirated on met mast - fan working/clean | y
Snow depth sensor - check free of spider webs/debris/grass and area under is flat | y
BF5 (domed radiation sensor) indicator blue (not pink) | y
ICOS EC mast (date, time, issue must be reported same day if issue)
Check TRH aspirated on ICOS mast - fan working/clean | y
Visually check sonic transducers and IRGA inlet (clean not blocked) | y
WEEKLY CHECKS
CABIN 1 |  | CABIN 3
Air Con Ok (Cabin 1) - SHOULD BE OFF IF FREE AIR ON | y | Air Con Ok (Cabin 3) - SHOULD BE OFF IF FREE AIR ON | y
Free air system running (%) | na | Free air system running (%) | na
Manifold Fan Running | y | Check CABIN3 desktop PC is turned on. | y
Electricity meter reading | 270360
Blog post written |  | SOUTH BOARDWALK
Email sent out |  | Partisol (ECOC) running and sampling time ≈ ~9360mins | y
|  | Partisol (ECOC) cassette number & filter end time recorded | y
CABIN 2 |  | HM Partisol Sampling time = ~360 minutes? | y
Air Con Ok (Cabin 2) - SHOULD BE OFF IF FREE AIR ON | y | HM Partisol cassette number & filter end time recorded | y
Free air system running (%) | na
NORTH-WEST BOARDWALK
|  | Heavy Metals Water Exchanged | y
LESS REGULAR TASKS |  | Big week (Fortnightly)
Little week (Fortnightly) |  | TOMPS Sample Exchanged | y
GC Spreadsheet Completed |  | TOMPS data downloaded | y
Bins emptied, replaced and returned to office? |  | TOMPS flow & sample volume set (120 & 999) | y
UKEAP Bulk Rain Exchanged |  | TOMPS next sample start date set (proceeding Tues pm) | y
New wet-only field blank |  | PAH Filters Exchanged | y
TOMPS running and flow OK? |  | PAH Bulk Rain Exchanged | y
TOMPS flow & sample volume set (F = ~120 & V = ~160) |  | PAH Head Swapped / Rotated | rotated
TOMPS fan symbol spinning |  | PAH head closed! | y
Recorded all wildlife sightings | y | Recorded all wildlife sightings | y
|  | MARGA yellow SJAC filters changed | na
Monthly
DELTA/ALPHA Samples Exchanged |  | 4-weekly tasks
Main Scaffolding Tower Check |  | HM Partisol Filters Added/Returned
Mercury Rainfall Exchanged | y | Digitel ECOC Sample Added/Returned
COMMENTS
Bucket under cabin 1AC had about 1.5 litres of water in, emptied
Engineers at ECOC and Heavy metal digitels - both on prestart until the right times
NO DI water at site, bring next week
Very strong chicken manure pile in gavins farm.
James swapped out CNR4 sensor for calibration. Logger prog updated.
3 empty MARGA waste bottles at site.
Note on A10: Windows User: Check for windows updates in system settings
Note on A12: Windows User: Double click on icon in system tray, then at the bottom right of the Sophos screen click on .About'.
Note on A14: RACKPC0114: Run the cmd window (enter "cmd" in the windows search box) Call "ntpq -pn" (note the space before the hyphen). Record the offset and jitter of the 192.168.0.187 ICOS Smartflux2 GPS time server.
Note on D35: Windows User: Currently logging to: C:\Data\Phenocam
Note on D36: Windows User: S:\Phenocam\Backup_REM1
Note on A39: Windows User: black 'sd' card icon
Note on D55: Windows User: Use the thermometer on the cabinet wall
Note on D66: RACKPC0114: Check latest data in BOTTLES.dat file Chageover time should be should be 23:55
Note on D70: Windows User: Make a note on blog & weekly email if width of tape on RH spool is less than 1cm Email Ivan, (cc Nathan & Duncan) if width of tape on RH spool is ≤ 0.5cm.
Note on D75: Windows User: No red light Status 'OK' and mode 'sampling'. Flow is approx 16.7 LPM. Sampled volume should be a single digit number of m3 - 1 m3 per hour since 6am (GMT) on the Wednesday morning.
Note on D76: Windows User: Take back to Bush and store filter in -20 Freezer before the end of the day.
Note on D83: Windows User: check file date
Note on A86: Windows User: Clean if dirty with lint free cloth (record below exact time of cleaning and report to Karen on same day).
Note on D86: Windows User: Clean if dirty with lint free cloth (record below exact time of cleaning and report to Karen on same day).
Note on A88: Windows User: Vegetation needs clearing regularly during summer months - report timings same day to Karen
Note on A93: Windows User: from ground only
Note on A102: Windows User: Use first KWh reading after time and date.
Note on D105: Windows User: Physically check most recent filter/cassette number. Record filter end time: On screen press Menu>read out internal memory >filter data>show data records. Record date/time of most recent record on sheet - this is the filter end date/time. Press back/Home to get back to main screen.
Note on D107: Windows User: Physically check most recent filter/cassette number. Record filter end time: On screen press Menu>read out internal memory >filter data>show data records. Record date/time of most recent record on sheet - this is the filter end date/time. Press back/Home to get back to main screen.
Note on D114: Windows User: USB stick should be on bench by Rem1
Note on A118: RACKPC0114: TOMPS should be running (there should be an icon of a turning motor). F should be approx 120, and V should be approx 160 (assuming logging started at midday on previous day)

## QC
start_time | end_time | QC_variable_name | comment
|  | list of names drop down | LOTS OF MANURE!!!

## logger name vs variable name li
logger_name | sensor_name | QC_name